## Slide 1: Models with multiple traps and centers
2 independent traps
2 traps including one deep trap not sensitive to the optical stimulation
2 traps including one shallow trap for which thermoluminescence is possible
2 recombination centers including a non-radiative center

## Slide 2: Independent traps
We consider 2 traps (1 and 2) optically sensitive and independent → we have the traps concentrations n1 and n2 and the escape probabilities p1 = ¿d1-1 and p2 = ¿d2-1 → superposition principle →
Superposition principle can be applied to more than 2 traps

## Slide 3: Additional deep trap without optical stimulation
Additional deep trap without optical stimulation (trap 2) occurs for a stimulation with long wavelength and thus with insufficient energy to empty the deep traps→ the deep trap is a competitor to the recombination center
The OSL intensity can be written (with N2, n2 and An2, the concentration of available traps, of filled traps and the trapping probability for the deep trap) →
The intensity is reduced due to the trapping in the deep trap
The decrease is no more exponential

## Slide 4: Additional shallow trap with thermoluminescence
The last 2 terms combine to produce a long-lived and temperature-dependent component to the OSL decay curve
That may imply (as a function of the parameters) an initial increase of IOSL followed by a decrease → in such circumstances the terms « OSL decay curve » is inappropriate

## Slide 5: Second recombination center (non-radiative)
A second recombination center (2) may exist and be non-radiative (or if radiative, outside the detector sensitivity window)
Let be m1 and m2, the h+ concentrations at R1 and R2
As only the recombination e-- h+ at 1 is radiative →
Again decrease of IOSL

## Slide 6: Comparison between the ≠ models

## Slide 7: More complex models
In previously models → we add 1 deep trap or 1 shallow trap or 1 recombination center
For realistic materials → they contain multiple traps, deep or shallow, and multiple recombination centres (radiative or not) → model with all these elements →

## Slide 8: Results for this complex model
p = 1
p = 0.1
p = 0.01

## Slide 9: LM-OSL
Alternative method introduced in 1996 by Bulur → the intensity of the stimulation linearly increases
The OSL response is a series of peaks, each peak corresponding to the optical escape of charge from a given type trap → the traps with the largest photoionization cross section are empty first
The traps characterized by slow, medium or fast escape are more easily discriminate than for CW-OSL
Not very often used in practice due to a larger technical complexity than for CW-OSL but also to a larger mathematical complexity

## Slide 10: 1 trap – 1 recombination center: 1st order
First order → Quasi-equilibrium assumption and Negligible retrapping
We consider for ©(t) →
This equation has a Gaussian as solution →

## Slide 11: Characteristic of the first order
The OSL curve shape starts from 0 at t = 0
The curve shows a maximum at
For known ° → it is possible to obtain ¾
The maximum intensity of the curve is →
The integration of IOSL gives n0

## Slide 12: Examples of LM-OSL curves
° = 0.3 m-3s-2
° = 0.01 m-3s-2
° = 0.003 m-3s-2

## Slide 13: Multiple independent traps
The superposition principle can be applied → for k traps characterized by their occupation ni and their photoionisation cross section ¾i →
OSL curve → sum of k curves having each a maximum dependent on ¾i →
OSL response of quartz

## Slide 14: Characteristics at general-order
QE but important retrapping rate
For 1 type of trap →
Maximum tmax for a value of Imax →
Attention → for several types of traps → retrapping implies that the traps are not independent → the superposition is not possible as for 1st order → complexity
and

## Slide 15: POSL
POSL allows to efficiently separate luminescence and stimulation lights
Important decrease of the background due to the scattered stimulation light
Very powerful method for weak dose dosimetry

## Slide 16: Principles of POSL (1)
We consider several stimulation pulses (supplied by LED or laser) of different intensities ©i and durations Ti such that ©i Ti = ©T = constant (fixed absorbed energy by pulse).
By considering the 1st order model →
For a weak stimulation (¢n ¿ n → the concentration of charges released by pulse is negligible compared to the concentration of trapped charges) → ¢n / ©T → the charge released from the traps is approximately equal after each pulse

## Slide 17: Principles of POSL (2)
Normally when writing rate equations → we consider that each recombination event leads instantly to a photon emission event
In reality → the recombination center is first excited (luminescence center) before relaxation by light emission → existence of a built-in delay ¿ (lifetime of the excited state)
If we have ne, the concentration of excited states →
One photon is emitted when the excited state relaxes →

## Slide 18: Principles of POSL (3)
If ¿ ¿ T → conditions of quasi-equilibrium → dne/dt ≈ 0 and IOSL ≈ -dm/dt as usual (as for CW-OSL)
If ¿ ≥ T → ≠ situation → a certain number of centers relaxe after the end of the stimulation pulse
If ¿ À T → most of the photons are emitted after the pulse (for Al2O3:C → 90% of the light emerges after the pulse if it is narrow enough)
For t < T → IOSL ↗ when t ↗ (optical stimulation of the traps)
For t > T → IOSL ↘ when t ↗ (relaxation with a time constant ¿)

## Slide 19: Schematic illustration of POSL
POSL after pulses with © = 103, 102 and 22 u.e./s → we have the corresponding T: T = 6.6, 66 and 300 ms (¿ = 100 ms was chosen)
The ratio of the areas after and during the pulse ↗ for T ↘
Ф=22

## Slide 20: Experimental illustration of POSL
POSL for Al2O3:C stimulated by laser during 0.1 ms
Ratio of the areas after and during the pulse for Al2O3:C

## Slide 21: Determination of the dose by POSL
The total integration of the POSL curve corresponding to 1 pulse is equal to the charges released from the traps and is thus / to n0
The integration of the curve after the end of the pulse is also / to the charges released from the traps and is thus also / to n0 but the proportionality coefficient changes as a function of the width T of the pulses
With ¢n ¿ n → only a few stimulation pulses are enough → to make several measurements is possible
POSL especially applied with Al2O3:C but very promising