## Slide 1: Another effect of irradiation on a MOSFET
Another effect of irradiation on a MOSFET → ↗ of the number of traps (several orders of magnitude) at the SiO2-Si interface (« interface » means on both sides → in the Si and in the SiO2)
Phenomenon not completely explained
Observation → ½ of the traps created by the irradiation are donors and ½ are acceptors → in the Si more traps (of 2 types) → whether for a nMOS or for a pMOS → more difficulties to obtain the channel → |VTH| ↗ (+ distortion of the characteristic)
For a pMOS → this effect adds up to ¢VTH (because ¢VTH < 0) because of h+ trapping
For a nMOS → this effect offsets ¢VTH
In practice → use of a pMOS

## Slide 2: Modification of ID - VGS (global effect)
nMOS pMOS

## Slide 3: Remarks on the use of a MOSFET
Lower detection limit « relatively high » → main use in radiotherapy and curietherapy
For « classical» available MOSFET dosimeters → lower limit ≈ 1 cGy
Investigations for an improvement of this lower limit → currently: ≈ 100 ¹Gy
Measurements are available at any moment → dose or dose rate
The ¢VTH shift caused by previous measurements must be known (because of the accumulation of the h+ in the insulator)
Limited life time of a MOSFET detector

## Slide 4: Temperature effect
Very sensitive to temperature (→ as for diodes) → apparition of a term
Between 20 °C and 40 °C → ≠ of 50 mV for ¢VTH → ≠ of 10 cGy
Generally → 1 °C → mV
D = 100 Gy

## Slide 5: Solution to the T problem: dual MOSFET
Dual MOSFET: 2 identical MOSFET, build on the same Si chip and operating at 2 ≠ gate voltages
During irradiation → each element gives a voltage shift → the ≠ between the 2 voltage shifts gives a signal / to the dose
For a dual MOSFET → variation of ≈ 0.015 mV/°C

## Slide 6: Dual MOSFET

## Slide 7: RADFET
To improve the sensitivity of the detector → several linked MOSFET → RADFET (RADiation-sensing Field-Effect Transistor)
Properties identical to the dual MOSFET but larger sensitivity: « theoretically » < ¹Gy → practically less good

## Slide 8: Response as a function of the dose: High doses
VGS = 0 V
VGS = 3 V
VGS = 6 V
VGS = 9 V
nMOS for 1 MeV °

## Slide 9: Response as a function of the dose: Small doses
Dual pMOS for 6 MeV °

## Slide 10: Energy response
Energy responses obtained in ≠ studies (6 MeV °) for ≠ MOSFET → in all cases → non-linearity for E < 100 keV
Wang et al. (2005)

## Slide 11: Effect of the cumulated dose
As for diodes → the irradiation introduces supplementary defects in the material → modification of the sensitivity as a function of the dose (for high doses)
≠ Behaviour for nMOS and pMOS ↔ because ≠ types of defects and thus of traps
pMOS
nMOS

## Slide 12: Comparison TLD-Diode-MOSFET

## Slide 13: Applications (1)
Mobile MOSFET system from Best Medical → wireless

## Slide 14: Applications (2)
OneDose MOSFET system from Sicel Technologies Inc. → wireless