## Slide 1: General theory of volume recombination in an ion chamber (Mie’s theory)
Theory valid for a uniform ionization rate in the chamber
Ion chamber with plane-parallel electrodes
Cylindrical ion chamber
Spherical ion chamber
General extrapolation

## Slide 2: Ion chamber with plane-parallel electrodes (1)
We consider an ion chamber with k electrodes (area S) separated by a distance d
The OX axis is ? to the electrodes and oriented from the positive electrode to the negative electrode → anode (+) is at x = 0 and cathode (−) at x = d
The electric field produced by the ions is negligible (at first approximation) compared to the applied field E0
As E0 is uniform, the drift velocities of positive (u+) and negative (u−) ions are constant
The ionization rate q (number of ionizations per unit of time and volume) is assumed to be constant and uniform
0
X
d
x

## Slide 3: Ion chamber with plane-parallel electrodes (2)
The quantity of + charges produced between 0 and x in the time interval dt is equal to the quantity of charges crossing the plane X = x in the same time interval → with n+(x) and n-(x) the bulk densities of charges at the distance x of the anode →

## Slide 4: Ion chamber with plane-parallel electrodes (3)
The recombination rate R’ (per unit of surface and time - m-2s-1) is:
With qd, the number of ionizations per unit of surface and time, we define the collection efficiency of charges f as the ratio between Qcoll, the collected charge at the electrodes and Q, the charge produced in the chamber by a constant radiation field →

## Slide 5: Ion chamber with plane-parallel electrodes (4)
The drift velocity of a charge u is linked to the applied electric field E by u = ¹E (with ¹, the mobility – m2V-1s-1)
In this case the applied potential U0 = Ed →
with m2 a constant of the gas (for air NCTP m2 = 6.47 £ 10-5 V2sm-1)
Calculation accounting for the electric field produced by the space charges (→ non-uniform E) gives →

## Slide 6: Cylindrical ion chamber (1)
Same arguments can be applied (with u and ¹ dependent on the position r) to a cylindrical ion chamber (with a and b, the intern and extern radii of the electrodes and L, the length of the cylinder):
r
r
anode with radius a
cathode with radius b

## Slide 7: Cylindrical ion chamber (2)
For cylindrical chamber →

## Slide 8: Cylindrical ion chamber (3)
The recombination rate R’’ (per unit of length and time - m-1s-1) is:

## Slide 9: Cylindrical ion chamber (4)
with

## Slide 10: Spherical ion chamber
Same calculation as for a cylinder → with » and κ values corresponding to a sphere →

## Slide 11: General extrapolation
Previous relations can be extrapolated and written as →
with c a constant depending on the detector geometry and on the nature of the gas
≠ values of 1/Qcoll as a function of U0→ extrapolation for 1/U02 = 0 → 1/Q correct for a continuous irradiation

## Slide 12: Columnar recombination
If volume recombinations are negligible and only columnar recombinations have to be considered (particles with high ionization density or high pressure gas) → empirical demonstration of Jaffe-Zanstra →

## Slide 13: Recombination for pulsed radiation
We choose dt: width of radiation pulses and T: charges collection time
If a lot of pulses during T or dt > T → theory of continuous irradiation
If dt ¿ T (dt ≈ ¹s and T ≈ ms) → Boag’s theory
dt
T

## Slide 14: Boag’s theory for pulse irradiation (1)
w0: initial width of the irradiated area → evolution of w with time t: w(t) = w0 - (u+ + u-)t with w(T) = 0 → T = w0 /(u+ + u-)
Hypothesis: n+ = n- = n →
0
X
d
w0
+
-
u+
u-

## Slide 15: Boag’s theory for pulse irradiation (2)
Integration of the recombination rate during time T → RT: recombination rate per unit of surface - m-2) →

## Slide 16: Boag’s theory for pulse irradiation (3)
The collection efficiency of charges f →
If w0 = d →
If small ® →

## Slide 17: Boag’s theory for pulse irradiation (4)
For a cylindrical chamber → d must be replaced by
For a spherical chamber → d must be replaced by

## Slide 18: Atmospheric corrections: Air density (1)
The collected charge depends on the type and the mass of gas inside the chamber
Most of the time → chamber open to ambient atmosphere → equilibrium of the gas with environment → density of the gas (with [T] = K and [P] = Pa) →
Pw is the water vapor partial pressure in air
Pw depends the relative humidity (RH) of air defined as the ratio of water vapor partial pressure in air on the saturation vapor pressure at the same T → measure of the link between the content of water vapor in air and its maximal capacity in the same conditions

## Slide 19: Atmospheric corrections: Air density (2)
Examples of Pw:
  at T = 295.15 K and RH = 100% → Pw = 2643.33 Pa
  at T = 295.15 K and RH = 50% → Pw = 1321.73 Pa
Examples of modifications of the density ½: If P = 101325 Pa and Pw = 1321.73 Pa → ½h (humid air) < by 0.5% compared to ½a of dry air (Pw = 0 Pa)
All these parameters have to be frequently measured during a long series of measurements
The humidity effect is often neglected because it also implies a correction on W/e varying in the opposite direction with the density correction → net error extremely small (< 0.3 %)

## Slide 20: Atmospheric corrections: Effect of humidity on W/e (1)
For dry air →
For humid air (W/e)h < (W/e)a but the ½ of air ↘ for humid air and as
→ the effects counterbalance
However humidity of air also influences on dE/dx

## Slide 21: Atmospheric corrections: Effect of humidity on W/e (2)
We consider a cavity ionization chamber under a flux of X-rays (for instance) constant, © = c
We assume that it is filled successively with dry air sec and with humid air →
1st term > 1 and 2nd term < 1 → Qh/Qa is relatively « flat » → 1.0028 § 0.0003 (15% < RH < 75%)

## Slide 22: Effects of humidity